IVD

ToRCH IgG Plus Test System

Rx Only

REF

A81101G

| Institute Name | Date |
| --- | --- |

PRINCIPLE OF THE ASSAY

The ZEUS AtheNA Multi-Lyte ToRCH IgG Plus Test System is a multiplexed assay designed to detect IgG antibodies specific for Toxoplasma gondii, Rubella, CMV, HSV-1 gG-1 and HSV-2 gG-2 in human sera. The test procedure involves two incubation steps:

- Test sera (properly diluted) are incubated in a vessel containing a multiplexed mixture Bead Suspension.  The Bead Suspension contains a mixture of distinguishable sets of polystyrene microspheres (beads). Conjugated to the primary sets of microspheres are Toxoplasma, Rubella, CMV, HSV-1 and HSV-2 antigens. If present in patient sera, specific antibodies will bind to the immobilized antigen on one or more of the bead sets.  The beads are rinsed to remove non-reactive serum proteins.

- Phycoerythrin-conjugated goat anti-human IgG is added to the vessel and the plate is incubated. The Conjugate will react with IgG antibody immobilized on the solid phase in step 1.  The Bead Suspension is then analyzed by the AtheNA Multi-Lyte instrument.  The bead set(s) are sorted (identified) and the amount of reporter molecule (PE conjugate) is determined for each bead set.  Using the Intra-Well Calibration Technology®, internal calibration bead sets are used to convert raw fluorescence into outcome (units).

TEST SYSTEM COMPONENTS

Materials Provided:

Each Test System contains the following components in sufficient quantities to perform the number of tests indicated on the packaging label. NOTE: The following components contain Sodium Azide as a preservative at a concentration of <0.1% (w/v): Bead Suspension, Controls, Conjugate and SAVe Diluent®.

| SOLN | BEAD |  |  |  | 1. | Bead Suspension: Contains separate distinguishable 5.6 micron polystyrene beads that are conjugated with the following antigens: Toxoplasma, Rubella, CMV, HSV-1 and HSV-2. The Bead Suspension also contains one bead set designed to detect non-specific antibodies in the patient sample (if present) and four separate bead sets used for assay calibration. One, amber bottle containing 5.5mL. Ready to use. |
| --- | --- | --- | --- | --- | --- | --- |
| CONJ |  |  |  |  | 2. | Conjugate: Phycoerythrin conjugated goat anti-human IgG ( chain specific). One, amber bottle containing 15mL. Ready to use. |
| CONTROL |  | + |  | 1 | 3. | Positive Control 1 (Human Serum): One, red-capped vial containing 0.2mL. |
| CONTROL |  | + |  | 2 | 4. | Positive Control 2 (Human Serum): One, white-capped vial containing 0.2mL. |
| CONTROL |  | - |  |  | 5. | Negative Control (Human Serum): One, green-capped vial containing 0.2mL.. |
| DIL |  |  | SPE |  | 6. | SAVe Diluent®: One, green-capped bottle containing 50mL of phosphate-buffered-saline. Ready to use. NOTE: The SAVe Diluent® will change color when combined with serum. |
| WASHBUF |  |  | 10X |  | 7. | Wash Buffer Concentrate (10X): Dilute 1 part concentrate + 9 parts deionized or distilled water. One, clear-capped bottle containing containing 50mL of 10X concentrated phosphate-buffered-saline. |

NOTES:

- The following components are not Test System Lot Number dependent and may be used interchangeably with the ZEUS AtheNA Multi-Lyte Test Systems: Wash Buffer and SAVe Diluent®.

- Test System also contains:

  - Component Label containing lot specific information inside the Test System box.

  - Calibration CD containing lot specific kit calibration values required for specimen analysis and assay quality control, and Package Inserts.

  - One 96-well dilution plate.

  - One 96-well filter plate.

PRECAUTIONS

- For In Vitro diagnostic use.

- Follow normal precautions exercised in handling laboratory reagents. In case of contact with eyes, rinse immediately with plenty of water and seek medical advice. Wear suitable protective clothing, gloves, and eye/face protection. Do not breathe vapor. Dispose of waste observing all local, state, and federal laws.

- The AtheNA Multi-Lyte Bead Suspension does not contain viable organisms. However, the reagent should be considered potentially biohazardous materials and handled accordingly.

- The Controls are potentially biohazardous materials. Source materials from which these products were derived were found negative for HIV-1 antigen, HBsAg and for antibodies against HCV and HIV by approved test methods. However, since no test method can offer complete assurance that infectious agents are absent, handle these products at the Biosafety Level 2 as recommended for any potentially infectious human serum or blood specimen in the Centers for Disease Control/National Institutes of Health manual “Biosafety in Microbiological and Biomedical Laboratories”: Current Edition; and OSHA’s Standard for Bloodborne Pathogens (36).

- Adherence to the specified time and temperature of incubations is essential for accurate results. All reagents must be allowed to reach room temperature (20 - 25C) before starting the assay. Return unused reagents to refrigerated temperature immediately after use.

- Improper washing could cause false positive or false negative results. Be sure to minimize the amount of any residual wash solution; (e.g., by blotting or aspiration) before adding Conjugate. Do not allow the wells to dry out between incubations.

- The SAVe Diluent®, Bead Suspension, Controls, and Conjugate contain Sodium Azide at a concentration of <0.1% (w/v). Sodium Azide has been reported to form lead or copper azides in laboratory plumbing which may cause explosions on hammering. To prevent, rinse sink thoroughly with water after disposing of solution containing Sodium Azide.

- The Wash Buffer concentrate is an IRRITANT. It is irritating to eyes, respiratory system and skin.

- Dilution or adulteration of these reagents may generate erroneous results.

- Do not use reagents from other sources or manufacturers.

- Never pipette by mouth. Avoid contact of reagents and patient specimens with skin and mucous membranes.

- Avoid microbial contamination of reagents. Incorrect results may occur.

- Cross contamination of reagents and/or samples could cause erroneous results.

- Avoid splashing or generation of aerosols.

- Do not expose reagents to strong light during storage or incubation. The Bead Suspension and Conjugate are light sensitive reagents.  Both have been packaged in light protective containers.  Normal exposures experienced during the course of performing the assay will not affect assay performance.  Do not expose these reagents to strong sources of visible light unnecessarily.

- Collect the wash solution in a disposal basin. Treat the waste solution with disinfectant (i.e.: 10% household bleach - 0.5% Sodium Hypochlorite). Avoid exposure of reagents to bleach fumes.

- Caution: Neutralize any liquid waste at an acidic pH before adding to a bleach solution.

- Do not allow the Conjugate to come in contact with containers or instruments that may have previously contained a solution utilizing Sodium Azide as a preservative.  Residual amounts of Sodium Azide may destroy the Conjugate’s enzymatic activity.

- Do not expose any of the reactive reagents to bleach-containing solutions or to any strong odors from bleach-containing solutions. Trace amounts of bleach (sodium hypochlorite) may destroy the biological activity of many of the reactive reagents within this Test System.

MATERIALS REQUIRED BUT NOT PROVIDED

- Pipettes capable of accurately delivering 10 - 200µL.

- Multichannel pipette capable of accurately delivering 10 - 200µL.

- Reagent reservoirs for multichannel pipettes.

- Serological pipettes.

- Disposable pipette tips.

- Paper towels.

- Laboratory timer to monitor incubation steps.

- Disposal basin and disinfectant (i.e.: 10% household bleach - 0.5% Sodium Hypochlorite).

- AtheNA Multi-Lyte System (Luminex® Instrument) with Sheath Fluid (Product Number 40-50035).

- Distilled or deionized water.

- Vortex.

- Small Bath Sonicator.

- Plate shaker capable of shaking at 800 RPM (optional for mixing).

- Vacuum aspirator and vacuum manifold for washing the microspheres.

STORAGE CONDITIONS

|  | Bead Suspension: Remove only the required amount to analyze the specimens to be tested and return the unused portion to storage. |
| --- | --- |
|  | Conjugate: DO NOT FREEZE. |
|  | Unopened Test System, Positive Controls, Negative Control, SAVe Diluent® |
|  | Wash Buffer (1X): 20 - 25C for up to 7 days, 2 - 8°C for 30 days. Wash Buffer (10X): 2 - 25C |

SPECIMEN COLLECTION

- ZEUS Scientific recommends that the user carry out specimen collection in accordance with CLSI document M29: Protection of Laboratory Workers from Infectious Disease (Current Edition).

- No known test method can offer complete assurance that human blood samples will not transmit infection. Therefore, consider all blood derivatives potentially infectious.

- Use only freshly drawn and properly refrigerated sera obtained by approved aseptic venipuncture procedures in this assay (13, 14). Do not use if there are any added anticoagulants or preservatives.  Avoid using hemolyzed, lipemic, or bacterially contaminated sera.

- Store sample at room temperature for no longer than 8 hours. If testing is not performed within 8 hours, sera may be stored between 2 - 8°C, for no longer than 48 hours. If a delay in testing is anticipated, store test sera  at –20°C or lower. Avoid multiple freeze/thaw cycles which  may cause loss of antibody activity and give erroneous results. It is the responsibility of the individual laboratory to use all available references and/or its own studies to determine stability criteria for its laboratory (40).

ASSAY PROCEDURE

- Remove the individual components from storage and allow them to warm to room temperature (20 - 25C).

- Determine the total number of Controls and samples to be tested. It is necessary to include the Negative Control and the two Positive Controls with each run.  The Negative Control should be tested in well A1, Positive Control 1 in well B1 and Positive Control 2 in well C1.  Each Control and sample requires one microwell for processing.

  - To optimize read times, the Bead Suspension must be thoroughly mixed just prior to use.  The most effective for re-suspension is to first vortex for approximately 30 seconds followed by sonication for approximately 30 seconds in a small bath sonicator.

  - For proper performance, it is important that the contents of the assay are thoroughly mixed.  Suitable means of mixing include mixing the plate on a plate shaker for approximately 30 seconds at approximately 800 RPMs or to set a pipettor to roughly ½  of the volume in the plate and repeatedly aspirate and expel (pump up and down) the contents of the well for a minimum of 5 cycles.

| EXAMPLE PLATE SET-UP |  |  |
| --- | --- | --- |
|  | 1 | 2 |
| A | Negative Control | Etc. |
| B | Positive Control 1 |  |
| C | Positive Control 2 |  |
| D | Patient 1 |  |
| E | Patient 2 |  |
| F | Patient 3 |  |
| G | Patient 4 |  |
| H | Patient 5 |  |

- Prepare a 1:21 dilution (e.g.: 10µL of serum + 200µL of SAVe Diluent®) of the Negative Control, Positive Controls, and each patient serum. NOTE: The SAVe Diluent® will undergo a color change confirming that the specimen has been combined with the diluent. For proper performance, it is important that the sample dilutions are thoroughly mixed according to 2b above.

- After determining the total number of wells to process, use a multichannel or a repeating pipette to dispense 50µL of the Bead Suspension into each of the wells of the filtration plate.

- Transfer 10μL of each diluted sample (1:21) and Control from the dilution plate to the filtration plate. For proper performance, it is important that the sample dilution and Bead Suspension are thoroughly mixed according to 2b above.

- Incubate the plate at room temperature (20 - 25°C) for 30 ± 10 minutes.

- After the incubation, rinse the Beads by vacuum filtration using the supplied Wash Buffer diluted to the 1X concentration.

- Place the filtration plate on the vacuum manifold and remove the solution, leaving the beads behind.

- Turn off the vacuum and add 200µL of  1X Wash Buffer.

- Apply the vacuum and remove the solution.

- Repeat steps 7b and 7c for a total of three rinses.

- Following the final wash, gently blot the bottom of the filter plate and allow the plate to air dry for 3 - 5 minutes before proceeding to the next step.

- Add 150µL of the Conjugate to each well, at the same rate and same order as the specimens.  For proper performance, it is important that the Conjugate and Bead Suspension are thoroughly mixed according to 2b above. As an option, while mixing the Conjugate one may transfer the mixture to empty wells of a polystyrene reaction plate.

- Incubate the plate at room temperature (20 - 25°C) for 30 ± 10 minutes.

- Set the AtheNA Multi-Lyte instrument to analyze the reactions by selecting the ToRCH IgG Plus template.  Refer to the operators manual for details regarding the operation of the AtheNA Multi-Lyte instrument.  Results may be read from the filter plate or a reaction plate. NOTE:  For proper specimen analysis, it is important that the instrument is set-up, calibrated and maintained according to the manufacturer’s instructions.  Please review the instrument manual for instrument preparation prior to reading the assay results.

- The plate should be read within 60 minutes after the completion of the Conjugate incubation.  One may decide to shake the plate for approximately 15 seconds prior to reading.  This optional step may reduce the amount of time required to read the plate.

| ##### Step | Abbreviated Assay Procedure |
| --- | --- |
| 1 | Dilute specimens 1:21 in SAVe Diluent®. Mix well. |
| 2 | Combine 50µL of Bead Suspension and 10µL of diluted specimen in an empty well. Mix well. |
| 3 | Incubate at room temperature for 30 ± 10 minutes. |
| 4 | Rinse the microspheres 3 times with 200µL of 1X Wash Buffer. |
| 5 | Gently blot the bottom of the plate and air dry for 3 - 5 minutes. |
| 6 | Add 150µL of Conjugate to each well. Mix well. |
| 7 | Transfer to a reaction plate (optional). |
| 8 | Incubate at room temperature for 30 ± 10 minutes |
| 9 | Shake plate (optional). |
| 10 | Read results within 60 minutes. |

QUALITY CONTROL

- Each time the assay is run it is necessary to include the Negative Control (in well A1) and the two Positive Controls (in wells B1 through C1).

- Run validity is determined through the performance of the Positive and Negative Controls.  These criteria are analyzed automatically through Intra-Well Calibration Technology.

  - The Negative Control and the two Positive Controls must all be negative on the non-specific or control antigen bead.

  - The Negative Control must be negative for each and every analyte included in the Bead Suspension.

  - Each Positive Control must be positive for a predetermined group of analytes included in the Bead Suspension.  These ranges are encoded within the Calibration CD. Positive Control ranges may be viewed by clicking on the “Control Graphs” button of the AtheNA software and then clicking “Control Upper/Lower Limits.”

  - Failure to meet any of the above criteria must result in non-reporting of patient results. Consider the entire run invalid and repeat.

- Specimen validity is based upon the characteristics of the calibration beads and their interactions with the patient sera.  There are various parameters monitored automatically through Intra-Well Calibration Technology.  If any of the criteria are found to be out of specification, the patient’s results are considered invalid and should be repeated.  Should this occur, the data report will indicate the particular specimen which has been invalidated as well as a trouble shooting code.

- Additional Controls may be tested according to guidelines or requirements of local, state, and/or federal regulations or accrediting organizations.  External Controls must be representative of normal human serum since AtheNA Multi-Lyte’s calibration system is partially based upon the characteristics of the serum sample.  If the specimen formulation is artificial (not human serum), erroneous results may occur.

- Good laboratory practice recommends the use of Positive and Negative Controls to assure functionality of reagents and proper performance of the assay procedure.  Quality control requirements must be performed in conformance with local, state and/or federal regulations or accreditation requirements and your laboratory’ s standard Quality Control procedures.  ZEUS recommends that the user refer to CLSI EP12-A and 42 CFR 493.1256 for guidance on appropriate QC practices.

INTERPRETATION OF RESULTS

##### Calculations

##### Assay Calibration: The ZEUS AtheNA Multi-Lyte ToRCH IgG Plus Test System utilizes Intra-Well Calibration Technology.  Intra-Well Calibration Technology includes a multi-point standard curve within the Bead Suspension. With Intra-Well Calibration Technology, each well of the assay is calibrated internally without any user intervention.  The standard curve is designed to self-adjust based upon the unique characteristics of the patient or Control serum.  Calibrator values are assigned to the internal standards by ZEUS, are lot specific and are encoded within the lot specific Calibration CD.

##### Analyte Cutoff Values: Each analyte of the ZEUS AtheNA Multi-Lyte ToRCH IgG Plus Test System has an assigned cutoff value.  Cutoff values are determined by ZEUS for each test system lot, and are encoded within the lot specific Calibration CD.

##### Through Intra-Well Calibration Technology, all calculations are performed automatically when using the AtheNA Multi-Lyte system.  Intra-Well Calibration Technology performs a regression analysis of the internal standards and then adjusts the calculated unit values based upon an additional standard and the characteristics of the serum sample.

##### Interpretations

##### Result Interpretation For Toxoplasma, CMV, HSV-1 and HSV-2:

<100 AU/mL: Interpret result as negative. No significant amount of IgG antibody to Toxoplasma, CMV, HSV-1 or HSV-2 IgG detected.

>120 AU/mL: Interpret result as positive. IgG antibodies to Toxoplasma, CMV, HSV-1 or HSV-2 detected.

100 - 120 AU/mL: Interpret result as equivocal for presence of IgG antibody to Toxoplasma, CMV, HSV-1 or HSV-2. Re-test equivocal samples in duplicate. If upon re-testing, one of the two samples remains equivocal, test the sample by an alternate serological procedure such as ELISA or draw a second sample one to three weeks later and test.

INV NSC: Indicates too much activity on the non-specific control bead. Re-test the sample, if result remains invalid, re-evaluate with a fresh sample.

    - Result Interpretation for Rubella:

0 - 9 Interpret result as negative. Detection of IgG antibodies for Rubella did not occur. Presume that patient is not immune against infection with Rubella.

11 or Greater: Interpret result as positive. IgG antibodies to Rubella were detected at a level that indicates immunity to infection.

10: Interpret result as indeterminate for presence of IgG antibody to Rubella virus. Re-test indeterminate samples in duplicate. If upon re-testing one of the two samples remain equivocal, test the sample by an alternate serological procedure such as ELISA, or draw a second sample one to three weeks later and test.

INV NSC: Indicates too much activity on the non-specific control bead. Re-test the sample, if result remains invalid, re-evaluate with a fresh sample.

NOTE: ZEUS Scientific technicians calibrated this test to the first WHO International Standard for Rubella IgG at the cut-off.  The result of 11 is equivalent to the cutoff of 10 IU/ml of the WHO Standard for Rubella. The magnitude of the test result above or below the cut-off does not correspond to International Units and is not indicative of total amount of antibody present.

LIMITATIONS OF THE ASSAY

- Performance characteristics of this device have not been established for matrices other than serum.

- Samples collected too early in the course of the infection may not have detectable levels of IgG antibodies. Negative results may be due to delayed seroconversion.

- The numeric value of the result above the cutoff is not indicative of the amount of anti-Toxoplasma, Rubella, CMV, HSV-1 or HSV-2 IgG antibody present.

- False positive results may occur.  Repeat testing, or testing with a different device may be necessary in some settings e.g. symptomatic patients.

- The ZEUS AtheNA Multi-Lyte ToRCH IgG Plus Test System is a diagnostic aid and by itself is not diagnostic. Test results should be interpreted in conjunction with the clinical history, epidemiological data and other information available to the attending physician in evaluating the patient and the results of other diagnostic procedures.

- Hemolytic, icteric, or lipemic samples may interfere with the outcome of this assay.  Additionally, specimens with abnormal IgG concentrations may interfere with the outcome of this assay.  Use of these types of specimens should be avoided.

REFERENCES

- Krick JA, and Remington JS: Toxoplasmosis in the adult: An overview. New Engl. J. Med. 298:550-553, 1978.Anderson SE and Remington JS: The diagnosis of toxoplasmosis. So. Med. J. 68:1433-1443, 1975.

- Anderson SE, Remington JS: The diagnosis of toxoplasmosis. So. Med. J. 68:1433-1443, 1975.

- Welch PC, Masur H, Jones TC, and Remington JS: Serologic diagnosis of acute lymphadenopathic toxoplasmosis. J. Infect. Dis. 142:256-264, 1980.

- Wilson CB, Remington JS, Stagno S, and Reynolds DW: Development of adverse sequelae in children born with subclinical congenital Toxoplasma infection. Pediatrics 66:767-774, 1980.

- Hughes HPA: Toxoplasmosis: The need for improved diagnostic techniques and accurate risk assessment. Contem. Topics Micro. Immunol. 120:105-139, 1985.

- Lamoril J, Molina JM, et al: Detection by PCR of Toxoplasma gondii in blood in the diagnosis of cerebrospinal toxoplasmosis in patients with AIDS.J Clin Pathol. 1996 Jan; 49 (1): 89-92.

- Cooper LZ: The history and medical consequences of rubella. Rev. Infect. Dis. 7:502-510, 1985.

- Assad F and Ljungars-Esteves K: Rubella - world impact. Rev. Infect. Dis. 7:529-536, 1985.

- Norrby, E.: Rubella Virus, in: Manual of Clinical Microbiology, 4th ed., Lennette, E.H., A Balows, WJ Hausler, Jr., H.J Shadomy (eds).  American Society for Microbiology, Washington, DC. Pp: 774-778, 1985.

- Cooper LZ, Green RH, Krugman S, Giles JP, and Mirick GS: Neonatal thrombocytopenic purpura and other manifestations of rubella contracted in utero. Am. J. Dis. Child. 110:416-427, 1965.

- Miller E, Cradock-Watson JE, and Pollock TM: Consequences of confirmed maternal rubella at successive stages of pregnancy. Lancet 2:781-784, 1982.

- Hinman AR, Bart KJ, Orenstein WA, and Preblud SR: Rational strategy for rubella vaccination. Lancet 1:39-41, 1983.

- Herrman KL: Rubella virus In: Lennette EH, Balows A, Hausler WJ, and Shadomy HJ, eds., Manual of Clinical Microbiology, American Society for Microbiology, Washington, DC, Ch. 76, pp. 779-784, 1985.

- Jordon MC: Latent infection and the elusive cytomegalovirus. Rev. Infect. Dis. 5:205-215, 1983.

- Starr SE: Cytomegalovirus. Ped. Clin. N. Am. 26:282-293, 1979.

- Weller TH: Clinical spectrum of cytomegalovirus infection. In: Nahmias AJ, Dowdle WR, and Schinazi RF, eds. The Human Herpes Viruses, an interdisciplinary perspective. Elsevier/North Holland Publishing Co., New York, pp. 20-30, 1980.

- Adler SP: Transfusion-associated cytomegalovirus infections. Rev. Infect. Dis. 5:977-993, 1983.

- Reynolds DW, Stagno S, and Alford CA: Laboratory diagnosis of cytomegalovirus infections. In: Lennette EH and Schmidt NJ, eds. Diagnostic Procedures for Viral, rickettsial, and Chlamydial Infections, 5th ed., American Public Health Association, Washington, DC, 1979.

- Nankervis G: Long term follow-up of cytomegalic inclusion disease of infancy, Pediatrics 46:403-410, 1970.

- Stagno S, Reynolds DW, Huang ES, Thames SD, Smith RJ, and Alford CA: Congenital cytomegalovirus infections: Occurrence in an immune population. N. Engl. J. Med. 296:1254-1258, 1978.

- Handsfield HH, Chandler SH, Caine VA, Meyers JD, Corey L, Medeiros E, and McDougall JK: Cytomegalovirus infection in sex partners: Evidence for sexual transmission. J. Infect. Dis. 151:344-348, 1985.

- Jordon MC, Rousseau WE, Noble GR, Stewart JA, Noble CR, and Chin TDY: Spontaneous cytomegalovirus mononucleosis. Ann. Intern. Med. 79:153-160, 1973.

- Umetsu M, Chiba Y, Horino K, Chiba S, and Kakao T: Cytomegalovirus-mononucleosis in a newborn infant. Arch. Dis. Child. 50:396-398, 1975.

- Drew WL: Diagnosis of cytomegalovirus infection. Rev. Infect. Dis. 10:5468-5475, 1988.

- Nahmias AJ, and Roizman BR: Infection with Herpes Simplex viruses 1 and 2. New Eng. J. Med. 289:667, 719, 781, 1983.

- Lycke E, and Jeansson S: Herpesviridaie: Herpes Simplex virus. In: EH Lennett, P Halonen, and FA Murphy,eds., Laboratory Diagnosis of Infectious diseases: Principals and Practice, vol. II: Viral, rickettsial and Chlamydial Diseases, Springer-Verlag, Berlin, pp 211, 1988.

- Nahmais AJ, Dannenbarger J, Wickliffe C, and Muther J: Clinical aspects of infection with Herpes simplex viruses I and II. In: AJ Nahmias, WR Dowdle, and RF Schinzai, eds., The Human Herpes Viruses, an  Interdisciplinary Perspective, Elsevier/North-Holland Publishing co., New York, pp 2, 1980.

- Whitley RJ and Hutto C: Neonatal Herpes Simplex virus infections. Ped. Rev. 7:119, 1985.

- Brown Z, Sleke S, Zeh J, Kopelman J, Maslow A, Ashley R, Watts D, Berry S, Herd H, Correy L. The acquisition of herpes simplex virus during pregnancy. N Engl J Med 337:505-515 (1997)

- Aurelian, L.Herpes Simplex Viruses. 473-497. In Specter, S & G Lancz, eds. , Clinical Virology Manual. 2nd Ed. Elsevier, New York. (1992)

- Centers for Disease Control and Prevention. Sexually transmitted diseases treatment guidelines 2002. MMWR 2002:51 (No.RR-6)

- Arvin, A, C.Prober. Herpes Simplex Viruses. 876-883. In Murray, P, E. Barron, M. Pfaller. F. Tenover, and R. Yolkenet (eds.). Manual of Clinical Microbiology. 6th Ed. ASM, Washington, D.C. (1995).

- Munday, P.E., J. Vuddamalay, M.J. Slomka, and D.W.G. Brown. 1998. Role of type specific herpes simplex serology in the diagnosis and management of genital herpes. Sex. Trasm.Infect. 74: 175-178.

- Wald A, J Zeh, S Selke, et al. (2000) Reactivation of Genital Herpes Simplex Type 2. Infections in asymptomatic seropositive persons. N.Eng.J.Med 342:844-849.

- U.S. Department of Health and Human Services. Public Health Service. Centers for Disease Control and Prevention and National Institutes of Health.  U.S. Government Printing Office, Washington D.C., 4th Ed., 1999.

- U.S. Department of Labor, Occupational Safety and Health Administration; Occupational Exposure to Bloodborne Pathogens, Final Rule. Fed.Register 56:64175-64182, 1991.

- Protection of Laboratory Workers from Instrument Biohazards and Infectious Disease Transmitted by Blood, Body Fluids and Tissues; Approved Guideline.  NCCLS/CLSI Document M29, Vol.17 (12), 1997.

- CLSI H18-A2: Procedures for the Handling and processing of Blood Specimens, 2nd Ed. (2006)

- CLSI EP7-A2: Interference Testing in Clinical Chemistry; Approved Guideline, 2nd Ed. (2005).

- Procedures for the Handling and Processing of Blood Specimens for Common Laboratory Tests; Approved Guidelines – 4th Edition (2010).  CLSI Document GP44-A4 (ISBN 1-56238-724-3).  Clinical and Laboratory Standards Institute, 950 West Valley Road, Suite 2500, Wayne, PA 19087.

In the US & Canada, call toll free, or e-mail Customer Service (orders@zeusscientific.com) or Technical Service (support@zeusscientific.com).

For all other countries, please contact your local distributor.

©2021 ZEUS Scientific All Rights Reserved.

ZEUS Scientific

200 Evans Way, Branchburg, New Jersey, 08876, USA

Toll Free (U.S.): 1-800-286-2111, Option 2

International: +1 908-526-3744

Fax: +1 908-526-2058

Website: www.zeusscientific.com

AtheNA Multi-Lyte and SAVe Diluent® are trademarks of ZEUS Scientific